| Design Input | Design Output | Verification Method |
| --- | --- | --- |
| The system must classify dental risk into Low/Moderate/High | Questionnaire logic + ML scoring file | Unit testing of scoring; accuracy tests with synthetic data; timed usability test |
| The app must detect visual caries indicators | Trained CNN model | Clinical validation vs. labelled dataset (target Sens ≥0.85, Spec ≥0.80) |
| The app must operate on affordable smartphones | Device requirement spec | Performance testing on low-tier Android device |
| The app must minimize false reassurance in serious cases | Fusion logic with red-flag overrides | Branch coverage tests; scenario-based testing using extreme conditions |
| The app must protect user health information | Encryption & data schema documentation | Penetration tests; security audit; inspection of key management |
| The app must store as little personal data as possible | Privacy & consent flow | Manual verification + automated test confirming data deletion works |
| The app must handle poor images safely | Auto-quality-reject algorithm | Test with blurred/low-light images; acceptance rate ≤5% of poor images |
| The UI must be accessible & understandable | Wireframes + style guide | WCAG 2.1 AA checklist; readability analysis |
| The app must complete assessment within 10 min | Combined questionnaire + image capture outputs | User time trial with n≥10 participants |
| App results must be explainable | Triage module with reason codes | Inspection test: each output contains at least 1 justification item |
| App must operate safely under failure conditions | ISO 14971 Risk File + FMEA | Cross-check that each high-risk failure has mitigation implemented |
| App must be verifiable & traceable | Verification Plan | Instructor/internal review; mapping tests to outputs |

Comment:
- Design Input needs to match to Design Input in Chapter 1 doc

- Designed output is not specified enough -> need details

- Need to include the references in the other document -> need to know why it is >=0.85, need a way to store these documents

Goal for next step:

- Map Chapter 1 inputs to listed output

  - Likely a table of design input -> design output -> reason behind mapping

- Need to collapse design outputs together so its not messy

- After redefining outputs we can redo traceability martix